## Slide 1
Contents
Part 1 – Team, projects and infrastructure Team and infrastructure news
Part 2 – Instrumentation highlights Sensors, systems and methods R&D
Overview of Magnet Test Activities at CERN
M. Buzio on behalf of the Test & Measurement team

## Slide 2
Part I Team, Projects, Infrastructure

## Slide 3
The Team
SM18 (superconducting test station) and magnetic measurement teams merged in 2020
Now 45 strong (20 staff, 17 students, 8 associates and industrial support)
Slowed-down but not stopped by Covid 19 !

## Slide 4
High-Luminosity LHC
currently under test: MQXFB Q2 P3, D2 prototype
2026
https://espace.cern.ch/HiLumi
Notes: Highest-priority project: - replace rad-damaged IR quads - 10  luminosity = +25% discovery range
100 cryomagnets to be tested <2026
First operational accelerator-quality Nb3Sn magnets
11 T in stand-by, detraining with thermal cycles
Nb3Sn: no direct He c
ooing (no voids), operation at low field with large thermal margin -> large heat capacity -> slow quench + noise = difficult quench detection
Extend and “shuffle” cryogenic lines from CFB to HL-LHC magnets
Add a phase separator to control LHe level and temperature (4.5 K tests)
Add 400 W heaters to pre-load He control system (increased ramp losses)
Add instrumentation
Upgrade of vertical/horizontal test benches
Shuffling modules with phase separators
Innovations: PLC interlocks with SCADA …
New UQDS protection/DAQ system

## Slide 5
Carpenter test logbook
Oracle DB interface via OpenShift, developed in PHP on Laravel framework
Used for all cryomagnet tests in SM18 for safety, QA and result visualization
Levels and sequence of tests conform to fixed templates
Different user roles for each step
Each step linked to formally approved procedure documents and checklists
Standardized stats and plots include timeline of temperature, current etc.
Current test status and responsible engineers clearly marked
Credit: F. Mangiarotti
Status, date/time and operator in charge of each step appear in the test chronology

## Slide 6
Enhanced Nb3Sn racetrack model coil eRMS
16.5 T peak field @ 13.8 kA, 1.9 K
High Field Magnets for the Future Circular Collider
J C Perez, C. Petrone et al, “Construction and Test of the Enhanced Racetrack Model Coil”, IEEE Trans. on Appl. Supercond. Vol 32 n. 6, Sep 2022
5×transversal coil array PCB in the middle of the gap
3× LakeShore HGT 2101 Hall probes

## Slide 7
26 multiplets (ASG, 170 magnets), 192 mm warm bore, tested with rotating coil scanner/SSW
24 bending dipoles (Elytt), super-ferric with supercritical He cooling, 380  180 mm2 warm bore, tested with translating fluxmeter/SSW
3 reconfigurable benches ready in b. 181
Test campaign now ramping up …
GSI/FAIR - SuperFRS
Credit: P. Kosek
First main bending dipole being tested with translating fluxmeter
First short multiples being tested with SSW

## Slide 8
Part II Instrumentation Highlights

## Slide 9
600 mm long, interchangeable PCB design for bores  90 to 150 mm
Measurement of warm HL-LHC MQXFB magnets @ 10 A, 4.4 mT @ 50 mm
3 units at CERN + 1 deployed to CIEMAT
Warm rotating coil scanner
PCB-mounted retroreflector
hand-operated longitudinal positioning with prolongations
tilt sensor for gravity leveling
Credit: P. Rogacki, L. Fiscarelli
interchangeable PCB

## Slide 10
Long rotating coil shafts
4 × Ø109 mm, ~8 m long shafts being procured for HL-LHC magnets
Prototype unit w/ PCB coil array sandwiched between high-precision C-fiber half-shells
Good results, but price hikes and supply problems …
Double C-fiber half-shells for torsional and flexural stiffness 1.4 m long ~50 µm tolerance
6 micro-connections per segment
Credit: R. Beltron, O. Dunkel, G. Deferne, L. Fiscarelli
MQXFB integrated field
50-channel Mobile Rotation Unit
Shaft prolongation (no coils)
Sliding tracks
MQXFB harmonics hysteresis

## Slide 11
Ø110 mm Quench Antenna
12-segment quench antenna designed to improve quench localization in critical MQXF quads
New design: rolled-up flexi PCB
On-board analog bucking → B3, A3, B4, A4 outputs
Extremely easy to assemble/disassemble
Quench reconstruction resolution: - tangential: 5°
- radial: 5 mm
Bypass of analog-bucking op-amp on existing amplifier cards (G=100 to 500)
Zin adapted to compensate the capacitive load of long coax cables
10× S/N improvement !
B4
A4
B3
A3
2 × double-layer flexi PCB glued together, 4 × 16-turn tangential coils per layer
Ø100 mm, 0.6 long antenna segment
5 pins, 1 used, tape
Credit: J. Kaeske, P. Rogacki, V. Di Capua, R. Beltron, L. Fiscarelli

## Slide 12
Small-diameter rotating coils for INFN-LNF’s EuPRAXIA
5×256-turn, 32-layer radial PCB coils inserted and pinned in the C-fiber tube
Credit: A. Lauria, M. Pentella
Low-cost Ø26 mm C-fiber tube (20 µm RMS tolerance)
Ceramic ball bearings
Adjustable anodized Alu supports
Maxon DC motor
Accura 25® 3D-printed parts with 0.2 mm tolerance (required machining)
Harmonic torsional vibrations @ 4, 8, 15, 30, 90 Hz visible in the outermost coil signal (< 1%)
Vcoil output at 0.25 rev/s
<0.5 units RMS @ 10 mm low-order differences w.r.t. other instruments

## Slide 13
58 new/refurbished magnets for fixed-target East Area measured over ~one year
Extensive use of SSW for integral harmonics
Rotating DC mode SSW
0.2 units RMS difference
Credit: R. Chritin, C. Petrone, P. Rogacki
Ø140 mm quad used for cross-validation
up to 128 × 3 mm × 3 mm standard DC SSW measurements
Harmonics fitted to actual flux spanned
Notes: 64/128 points,
3 mm
Fit includes harmonics

## Slide 14
Pulsed excitation mode SSW
Create a 1-turn virtual coil by subtracting measurements at different wire positions
50-ppm excitation current reproducibility
Pulsed-mode flux integration in PS Booster main dipoles
Standard static SSW measurement added in to recover initial level at I0
C Petrone, J Vella Wallbank et. al, “Implementation of the SSW system in pulsed mode”, to be submitted to Sensors
Wire offset X2-X1: trade-off between higher S/N and effect of B3
200 ppm RMS difference vs. fixed integral coil

## Slide 15
AC excitation mode SSW
Goal: warm check of axis/field direction in HL-LHC corrector package MXBXFA
Stepwise circular wire movement to mimic a radial search coil:
vs. longitudinal position
A3 B3 A4 B4 A5 B5 A6 B6 A3
SSW Stage A
SSW Stage B
Magnets excited sequentially @ 19 Hz (compromise sensitivity/eddy current effects/available voltage)
N=1 central + 256 tangential 6-second acquisitions (oversampling)
return wire
Validated vs. DC-SSW on NC quad
Credit: E. Dalane, C. Petrone

## Slide 16
Used for warm measurement of HL-LHC correctors
Flux measurement taken on-the-fly
Translating Transversal Fluxmeter
Prototype Accura 25® 3D-printed head
12×30° 160-turn, 16-layer PCB coils
One additional coil shifted by 15° /30 mm to improve pitch/yaw misalignment
C-fiber moving rod and guiding tube (square !)
C-fiber guiding tube
Sick® PFG08 wire draw encoder (14 µm resolution, 300 µm reproducibility)
Longitudinal field profile of individual correctors in HL-LHC CP (deconvolved with coil sensitivity kernel) Localization accuracy << target 1 mm
Credit: E. Dalane, C. Petrone
Notes: Longitudinal position of in corr pack 1 mm required

## Slide 17
Translating Axial Fluxmeter
Goal: Quick in-situ measurement of field map/magnetic axis of 20 solenoids in the upcoming Antiproton Decelerator electron cooler
SSW not possible (clearances, magnet cross-talk)
Based on translating disc-shaped 8-layer PCB coil array
4×360° + 1 central, 6-turn concentric coils for axial and radial field profile
4×4×90°, 6-turn coils for magnetic axis localization
Axis defined by centers @ both end gradient peak planes 100 µm axis resolution with one pass 50 µm by repeating 9 passes w/ 5 mm offset
10 ppm field resolution
C Petrone, S. Russenschuck et al. , “Induction-coil measurement system for normal and superconducting solenoids”, IEEE Trans. on Appl. Supercond. Vol 32, n. 6 Sept 2022

## Slide 18
Real-time measurement systems
6 B-train measurement systems now operational in all normal-conducting CERN rings
250 kHz data rate, 30 µs latency, 2 µT resolution, ~100 µT absolute accuracy
Complex, many challenges, simplified systems for smaller scale applications under study
NMR probes
LEIR main bending dipole
Sensors in the fringe field region
coil
FMR marker
ISOLDE spectrometer challenges: only fringe region accessible, no spare, no validated FE model
Integral coil proposal for PS system (U17 unit under test in b. 867)
Simplest solution under test so far …
Credit: A. Beaumont, V. DI Capua, M. Pentella

## Slide 19
Arepoc HHP-NP 2067 Hall Probe
594 cm2 160-turn 16-layer PCB coil
Problem: indefinitely long integration of fixed-coil output
Kalman filtering for optimal estimation of the field in the presence of model (voltage offset V0) + measurement noise
Combining coil/Hall probe → three orders of magnitude improvement
Integrator drift cancellation with data fusion
V. Di Capua, M. Pentella et al., “Drift-free integration in magnetic measurements achieved by data fusion ”, Sensors 2022, 22, 182
Uncorrected drift 60 ppm/s
State-space model
Field = hidden state
Coil voltage = input variable
Case I: measurement = Hall probe
Case II: measurement = excitation current
V0 estimated 0—50 s 53 ppm/s
V0 updated on plateaus 3 ppm/s
Kalman + Hall probe 0.03 ppm/s
Kalman + exc. current 0.08 ppm/s
Actual eddy current decay transient
DCCT

## Slide 20
Thanks for your attention
… don’t miss these other presentations from CERN:
Improving the performance of a three-axes Hall probe mapper system (Melvin Liebsch)
“Hysteresis Modeling in Iron-Dominated Magnets Based on a Multi-Layered NARX Neural Network Approach” (Vincenzo Di Capua)
“A Static-Sample Magnetometer for Characterizing Weak Magnetic Materials”(Mariano Pentella)
“Encoderless harmonic coil measurements” (Piotr Rogacki)