## Slide 1
A Stretched Wire
Tomography System at Kyma
Mirko Kokole
… riding the wave

## Slide 2: Introduction
Technique of SW tomography first proposed and demonstrated by ESRF at IMMW19 in Grenoble in 2019
Kyma already has a Small Stretched Wire Bench
  for characterization of single magnets and small modules
Adaptation of the SW Bench to a SW Tomography system
  Addition of only one rotary stage necessary

## Slide 3: Radon Transform and Field Integrals
Radon transform is a path integral along the line L(,) of the function I(x,y).
Field integral is a path integral of the magnetic field density along a specific line.
Field integral over a line L is equivalent to the Radon transform of the magnetic field along the line L.
From field integral measurements scans at different angles we can reconstruct a 2D map of the magnetic field density.
A stretched wire bench with a rotary stage can be used to measure 2D magnetic field map.

## Slide 4: A Simple Back Projection - Laminograms

## Slide 5: Inverse Radon T. and Fourier Slice Theorem
Fourier transform of a projection function
A projection function
Inverse radon transform is a Fourier transform of a multiplied by a filter function .

## Slide 6: Filtered Back Projection

## Slide 7: Small Stretched Wire Bench
Basic design parameters
  Small size convenient to transport.
  Control by Raspberry Pi (SBC)
  Routines written in Python for ease of use
  Keithley Nanovolt 2182A meter

## Slide 8: Stretched Wire Bench - Construction
Fixed single strand wire loop
Linear translation stage
Rotary stage

## Slide 9: First Measurements
A set of 4 cylindrical magnets was used for the first measurements.

## Slide 10: Measurements of a magnet block
A segmented magnet block made from 3 glued pieces was used for tests.

## Slide 11: Axis shift correction
Jun K, Yoon S. Alignment Solution for CT Image Reconstruction using Fixed Point and Virtual Rotation Axis. Sci Rep. 2017 Jan 25;7:41218. doi: 10.1038/srep41218. PMID: 28120881; PMCID: PMC5264594. https://arxiv.org/ftp/arxiv/papers/1605/1605.04833.pdf
Each point in an image follows a sinus curve on a sinogram.
Misalignment of the rotation axis can be corrected by removing a sinus fit from a sinogram.

## Slide 12: Comparison with a Hall probe
SW Tomography
  Distance from surface was ~ 0.5 mm
  Field integrals measured at 180 angles
  Horizontal step was 0.4 mm/point
2D Hall probe field map
  Distance from surface 1.5 mm
  Longitudinal and horizontal step 0.5 mm/point
  128 x 128 point field map was created

## Slide 13: Conclusions
A dedicated SW Tomography bench was designed and assembled
First tests are very promising.
Advantages:
  Magnetic field very close to the surface can be measured.
  A dedicated bench can be used for magnet quality control
  Data could be used for further optimizations
Open questions:
  Can we measure absolute values of magnetic field?

## Slide 14: Thanks to the Team!
Special thanks to Andrej Nabergoj and Žan Mahnič