## Slide 1: A Static-Sample Magnetometer for Characterizing Weak Magnetic Materials
Mariano Pentella,
CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22)
Campinas, Brazil (Online)
29 September 2022
mariano.pentella@cern.ch

## Slide 2: Outline
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
Introduction
Magnetic characterization of materials
Characterization of weak magnetic materials
Static-sample magnetometer
Experimental results
Summary and conclusions

## Slide 3: Introduction
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
Materials have to display - thermal, electrical, mechanical, magnetic - properties compatible with the target application
In particle accelerator applications assessing how a material behaves under the effect of a magnetic field is crucial
Normal conducting and superconducting magnets
Magnetic shields
Collars, vacuum chambers, beam screens, etc.
In particular, many elements need to be “transparent” to the applied field to not perturb it, and thus not influence the beam stability.
Measurements of low magnetic permeabilities are required

## Slide 4: Magnetic characterization of materials - 1
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
The measurement methods can be grouped in three macro-areas
Fluxmetric methods: the sensing element, generally a pick-up coil, detects flux variations. The cause of the flux variations determines a sub-classification:
Application of a time-varying magnetic field .
Transformer principle
Magnetometer principle
Toroidal permeameters, bars, rods, Epstein frame, etc.
Vibrating and rotating sample magnetometers
The specimen is moved in a constant magnetic field .

## Slide 5: Magnetic characterization of materials - 2
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
Flux distortion methods: the presence of a test specimen having under a uniform field determines a local distortion of the field lines. This distortion can be mapped by a sensor and the permeability (or the magnetization) can be evaluated by inverse-problem solving.
A vast choice of sensor technologies can be used for this approach: fluxgate magnetometers, Hall probes, NMR
The measurement methods can be grouped in three macro-areas

## Slide 6: Magnetic characterization of materials - 3
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
The measurement methods can be grouped in three macro-areas
Force detection techniques (Gouy/Faraday balance): the force on a test specimen under the effect of an applied magnetic field can be detected by a change in the apparent mass. This apparent change of mass is proportional to the magnetic permeability of the specimen.

## Slide 7: Characterization of weak magnetic materials - 1
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
Flux-metric method using an open sample (brick shaped)
Magnetic permeability and repeatability bars measured on a specimen of stainless-steel type 316L
The lower the permeability, the more the uncertainty of the -field measurements has to decrease
Measured on a co-wound specimen of stainless steels for ITER voltage taps ()
The main challenge of measuring materials having is that we need to apply a strong field, such that
Toroidal permeameters commonly used for ferromagnetic steels do not have enough resolution for this kind of application
Even if we do that, however…

## Slide 8: Characterization of weak magnetic materials - 2
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
Stainless steel type 316 – VSM specimens vs SSM specimen (black)
Stainless steel type 316 after EDM cut – FIB-SEM scan
Vibrating sample magnetometry (VSM) represents the reference solution due to its resolution (). However, it presents two main limitations:
Given the typical specimen dimensions (of a few ), the mechanical preparation of the specimen can significantly impact the permeability to be measured and introduce a bias
Multiple samples are necessary to obtain a result which can be considered as representative of a material batch
Permeability map of a stainless-steel type 304L plate for the CMS HGCAL (High-Granularity CALorimeter). Measured at 0.1 T.

## Slide 9: Characterization of weak magnetic materials - 3
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
Bulky specimens can be measured effectively using flux-distortion detection techniques, which are reported in the ASTM and IEC standards. However, they have never been fully exploited in the literature so far.
Foerster Magnetoscope working principle
Portable instrument mainly used for in-situ inspections of finished parts
Limited accuracy (5 % of measured value)
No field sweep. Measurements possible only at 80 kA/m (about 0.1 T).
An important feature is that the instrument is not limited to a specific specimen shape
Can we improve the measurement accuracy, have the possibility to change the test field, without losing the flexibility in terms of specimen geometry?
Such measurement principle is generally adopted in commercial portable instruments, such as the Foerster Magnetoscop

## Slide 10: Static-sample magnetometer - 1
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
The model used for the Finite Element (FE) Analysis is a magnetostatic model in terms of reduced magnetic scalar potential formulation.
List of symbols
measured background field
measured perturbed field
relative permeability of the specimen under test
Boundary conditions set to 0
Example of output of the problem

## Slide 11: Static-sample magnetometer - 2
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
NMR Pt2025 teslameter, 0.1 resolution, 5 ppm absolute accuracy
Reference dipole, field range 0-1.2 T
Field map of the dipole magnet measured by NMR at an excitation current of 316 A.
PI linear positioning system
Integrated field quality (courtesy of M. Liebsch), measured by SSW

## Slide 12: Experimental results - 1
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
Cube
Interlab. check
10RM18 for HL-LHC magnets’ collars
Monte Carlo simulation results
WHA ASTM B-777 class 3 for radiation shielding (95 % tungsten)
Monte Carlo simulation results
JWO14159
Specimens:
A stack of laminations 60x60 mm, 3 mm thick
A stack of 6 toroid, 114mm De, 105mm Di, H 3 mm
Specimens:
12 bulk cylinders 35 mm diameter and 25 mm thick
Interlab. check

## Slide 13: Experimental results - 2
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
Characterization of AISI type 316L, grades 1.4404 and 1.4441
Specimens of raw material
Impact of deformation
DQW cryomodule layout
Impact of deformation
Finished parts
We can use a CAD drawing to model simple finished parts

## Slide 14: Experimental results - 3
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
In the absence of an available drawing, 3D scans can be used and imported in the FE software, thus fully exploiting the potentiality of using FE Analysis
Measurement of a hip prosthesis made of F-75. Computed permeability 1.00093±0.00003

## Slide 15: Summary and conclusions
Mariano Pentella, CERN, TE-MSC-TM
22nd International Magnetic Measurement Workshop (IMMW22) Campinas, Brazil (Online)
29 September 2022
A Static-Sample Magnetometer (SSM) was conceived and developed to meet an increasing demand of weak magnetic material measurements, to be used in a broad spectrum of projects and applications at CERN
The measurement is based on the flux-distortion detection principle, combined with Finite Element Analysis. The permeability resolution is in the order of 10-5, and the uncertainty of 10-4, with the method showing unprecedented performance in comparison with commercially available instruments.
Similarly to commercially available instruments, we can test specimens having various shape and dimensions (compatible with the magnet gap). In addition, test at several field values are possible.
A cold version with cold specimen in a warm magnet bore is currently under study
Pasquale Arpaia, Marco Buzio, Alessandro Parrella, Mariano Pentella, and Pedro M. Ramos. "A Static-Sample Magnetometer for Characterizing Weak Magnetic Materials." IEEE Transactions on Instrumentation and Measurement 70 (2020): 1-9.
Mariano Pentella "Characterization of magnetic materials at extreme ranges of field, temperature, and permeability”, Doctoral Thesis, https://iris.polito.it/handle/11583/2964790, 2022